WaterHeaterStation - whs.py

#!/usr/bin/env python3

import atexit

import signal

import sys

import time

import RPi.GPIO as GPIO

from max1238 import Max1238

# GPIO Pins

VALVE_PIN = 17

# MAX1238 channels

CH_HOT = 0

CH_COLD = 3 # Still Not Impl on hardware

CH_FLOW = 1

CH_AMBIENT = 4

# ADC config

ADC_VREF = 4.096

ADC_MAX = (2**12) - 1

R_SHUNT_OHMS = 120.0

# Temperature transmitters C)

T_MAX_C = 150.0

T_MIN_C = -50.0

# Flow transmitter (GPM)

Q_MAX_GPM = 10.0

Q_MIN_GPM = 0

# Fail-safe configuration

MAX_RUN_MINUTES = 5.0

MIN_FLOW_GPM = 0.05

LOW_FLOW_TIMEOUT_S = 20.0

PRINT_PERIOD_S = 0.5

# Init

GPIO.setmode(GPIO.BCM)

GPIO.setup(VALVE_PIN, GPIO.OUT, initial=GPIO.LOW)

adc = Max1238()

adc.setup_adc()

def _fail_safe_close():

try:

GPIO.output(VALVE_PIN, GPIO.LOW)

except Exception:

pass

atexit.register(_fail_safe_close)

def _signal_handler(signum, frame):

print(f"\n[!] Caught signal {signum}. Closing valve and exiting.")

_fail_safe_close()

sys.exit(1)

signal.signal(signal.SIGINT, _signal_handler)

signal.signal(signal.SIGTERM, _signal_handler)

def _raw_to_voltage(raw: int) -> float:

if raw is None:

return float("nan")

return (float(raw) / ADC_MAX) * ADC_VREF

def _volt_to_span(val_v: float, span_max: float, span_min: float) -> float:

if val_v != val_v:

return float("nan")

i_loop = val_v / R_SHUNT_OHMS

if i_loop < 0:

i_loop = 0.0

norm = (i_loop - 4e-3) / 16e-3

return (norm * (span_max - span_min)) + span_min

def read_voltage(channel: int) -> float:

raw = adc.read_single(channel)

return _raw_to_voltage(raw)

def read_amb_temps() -> float:

return read_voltage(CH_AMBIENT) / 10e-3

def read_externel_temps() -> dict:

return {

"Hot": _volt_to_span(read_voltage(CH_HOT), T_MAX_C, T_MIN_C),

"Cold": _volt_to_span(read_voltage(CH_COLD), T_MAX_C, T_MIN_C),

}

def read_flow_gpm() -> float:

return _volt_to_span(read_voltage(CH_FLOW), Q_MAX_GPM, Q_MIN_GPM)

def draw_water(target_vol_gal: float) -> float:

if target_vol_gal is None or target_vol_gal <= 0:

print("Invalid target volume.")

return 0.0

print(f"Target: {target_vol_gal:.3f} gal")

vol_gal = 0.0

start = time.monotonic()

last_log = start

low_flow_start = None

# Open valve

GPIO.output(VALVE_PIN, GPIO.HIGH)

try:

t_prev = time.monotonic()

while vol_gal < target_vol_gal:

now = time.monotonic()

dt = now - t_prev

t_prev = now

if (now - start) > (MAX_RUN_MINUTES * 60.0):

print("[!] Timeout reached. Stopping.")

break

temps = read_externel_temps()

amb_temps = read_amb_temps()

flow = read_flow_gpm()

if any(v != v for v in temps.values()) or (flow != flow):

print("[!] Sensor read error (NaN). Stopping.")

break

vol_gal += max(flow, 0.0) * (dt / 60.0)

if flow < MIN_FLOW_GPM:

if low_flow_start is None:

low_flow_start = now

elif (now - low_flow_start) >= LOW_FLOW_TIMEOUT_S:

print("[!] Low flow persists. Stopping.")

break

else:

low_flow_start = None

if (now - last_log) >= PRINT_PERIOD_S:

print(

f"T_hot={temps['Hot']:.1f} C  "

f"T_cold={temps['Cold']:.1f} C  "

f"T_amb={amb_temps:.1f} C  "

f"Flow={flow:.2f} gpm  "

f"Vol={vol_gal:.3f} gal"

)

last_log = now

time.sleep(0.05)

except Exception as e:

print(f"[!] Exception: {e}")

finally:

_fail_safe_close()

print(f"Volume drawn: {vol_gal:.3f} gal")

return vol_gal

if __name__ == "__main__":

try:

draw_water(10)

except KeyboardInterrupt:

print("\n[!] Aborted by user.")



Water_heaters_testings - water_draw_scripts - FMTest.py

import RPi.GPIO as GPIO

from time import time

GPIO.setmode(GPIO.BCM)

GPIO.setwarnings(False)

FMPIN = 6

VPIN = 17

GPIO.setup(FMPIN, GPIO.IN, GPIO.PUD_UP)

GPIO.setup(VPIN, GPIO.OUT, initial=GPIO.LOW)

GPIO.add_event_detect(FMPIN, GPIO.RISING)   #add rising edge detection

while True:

target = input('Enter desired water volume in gallons:')

target = float(target)

if target == 0:

GPIO.output(VPIN, GPIO.LOW) #close valve and quit if no water is requested

print('Exiting program.')

quit()

print ('Drawing %.2f gallon(s).' % target)

volume = 0

numPulses = 0

GPIO.output(VPIN, GPIO.HIGH)    #open valve

start_time = time()

while volume < target:  #target test volume in gallons

if GPIO.event_detected(FMPIN):

numPulses += 1

volume = float(numPulses) / 500  # wh_2 k = 500

#print ('Pulses: %1f' % numPulses)

#print ('Volume: %f' % volume)

run_time = time()

elapsed_time = run_time - start_time

if elapsed_time > 180:

print ('Timeout Error.')

break



Water_heaters_testings - water_draw_scripts - ValveControl.py

import RPi.GPIO as GPIO

GPIO.setmode(GPIO.BCM)

GPIO.setwarnings(False)

VPIN = 17

GPIO.setup(VPIN, GPIO.OUT, initial=GPIO.LOW)

while True:

x = input('1 to open, 2 to close: ')

if x == '1':

GPIO.output(VPIN, GPIO.HIGH)

print ('\n Valve Open \n')

elif x == '2':

GPIO.output(VPIN, GPIO.LOW)

print ('\n Valve Closed \n')

